### Contents

- Parts Name ……………………………………………………………………… 1

- Specification ……………………………………………………………………… 3

- Installation ………………………………………………………………………… 5

- Operation ………………………………………………………………………… 7

- Installation & Operation For Attachments …………………………… 10

- Maintenance …………………………………………………………………… 10

- Troubleshooting ……………………………………………………………… 11

- Outfits …………………………………………………………………………… 13

1. Parts Name

A1.Body   A2.Seidentopf binocular head  A3. Eyepiece  A4. Nosepiece  A5. Objective

A6. Mechanical Stage  A7. Clamp  A8. Condenser   A9. Notch for attachment  A10. Filter

A11.Condenser holder  A12. Handle   A13. Power Input  A14. Fuse  A15. Lamp Plate

B1. Head thumb screw  B2. Diopter adjuster  B3. Aperture diaphragm lever  B4.Condenser adjusting screw (left & right)  B5. Field diaphragm revolver  B6. Condenser thumb screw

B7. B8. Stage moving knobs  B9. Potentiometer  B10. Condenser focusing screw  B11.Power Switch  B12. Tension adjustment ring  B13. Coarse focusing knobs  B14. Fine focusing knobs

B15. Lamp plate thumb screw

2. Specification

2.1 Total magnifications

| Objectives Eyepiece | 4× | 10× | 20× | 40× | 100× |
| --- | --- | --- | --- | --- | --- |
| 10× | 40× | 100× | 200× | 400× | 1000× |
| 16× | 64× | 160× | 320× | 640× | 1600× |

2.2 Objectives (with eyepiece 10X)

| Objectives | N.A. | Objective Field (mm) | Resolving Power (μm) | Working Distance（mm） |
| --- | --- | --- | --- | --- |
|  |  | Field Numberφ20 |  | Achromatic objective | Plan objective |
| Finite | 4× | 0.10 | 5 | 2.8 | / | 16 |
|  | 10× | 0.25 | 2 | 1.1 | / | 2.1 |
|  | 40×(spring) | 0.65 | 0.5 | 0.42 | / | 0.63 |
|  | 100× (oil, spring) | 1.25 | 0.2 | 0.22 | / | 0.45 |
| Infinite | 4× | 0.10 | 5 | 2.8 | 9.5 | 6.73 |
|  | 10× | 0.25 | 2 | 1.1 | 1.7 | 4.19 |
|  | 20×(spring) | 0.40 | 1 | 0.69 | / | 2.14 |
|  | 40×(spring) | 0.66 | 0.5 | 0.42 | 0.39 | 0.45 |
|  | 100× (oil, spring) | 1.25 | 0.2 | 0.22 | 0.12 | 0.12 |

2.3 The other specification

2.3.1 Mechanical tube length: 160mm

2.3.2 Conjugate distance: Finite 195mm,

Infinite

2.3.3 Head: Seidentopf binocular (trinocular) 30°,

Interpupillary adjustable distance is 50-75mm,

Diopter adjustable range ±5,

Anti-fungal systems

2.3.4 Nosepiece: Quadplex or quinplex nosepiece

2.3.5 Mechanical Stage: Size 190mmX140mm

X-Y travel 55mmX78mm

2.3.6 Focusing systems: Coaxial coarse and fine focusing knobs,

Coarse stroke 25mm, fine division 2 μm,

Condenser up-down range 25mm

2.3.7 Condenser: Abbe condenser, N.A.1.25, Adjustable aperture,

Aperture center is adjustable

2.3.8 Illumination: Koehler illumination system

2.3.9 Filter: Blue filter is for standard outfits, green and amber is optional

2.3.10. Electric components: Input voltage AC 85-265V, 50/60Hz

Output voltage DC1.2-6V

6V/20W halogen lamp

Rotation potentiometer

Fuse 2A φ5×20

3. Installation

Please install the microscope as the following:

Installation Process:

- Installing Lamp (if the lamp is installed before purchase, please pass it)

a. Take down the base of microscope, Loose the screwB15, put out the lamp plate (Fig.3);

b. Take a new lamp from the lamp package, and hold the lamp with the foam;

c. Insert the lamp into lamp holder as the picture, then put back A15 lamp plate and the screw.

d. Take the microscope up and take the power.

- Installing binocular head

Loose B1 binocular head thumb screw, put out A2 head, put it on the top of body as Fig.1, tighten the screw.

- Installing objectives

Take out the objective from the packing box, and drive them into the holes of nosepiece orderly and tightly according to the times.

- Installing eyepiece

Take out the eyepiece tube cover, then get the eyepiece from the packing bag, and insert it into the tubes.

- Installing condenser

  - Turn B13 coarse focusing knob, rise A6 mechanical stage to top;

  - Turn B10 condenser focusing screw, and take A11 condenser holder to the lowest position;

  - Loose B6 condenser thumb screw, put the condenser to the holder, tighten B6 condenser thumb screw;

  - Turn B10 condenser focusing screw, and rise the condenser to working position.

- Installing blue filter or the other attachment

If blue filter would be used, the installing should be done before process 5, or do the installing during process 5-C, then put A10 blue filter or the other attachment to the bottom hole of A8 condenser.

- Installing the module attachment

The microscope can be with A9 module attachment, and it is multifunctional. Analyzer and wavelength plate can be installed in A9 notch. Polarizer is installed over field diaphragm or below the condenser. Filter can be installed in the notch if the other attachment is in the bottom hole of condenser.

- Installing CCD adaptor

Loose the thumb screw in trinocular tube, put the CCD adaptor into trinocular tube, then tighten the screw.

4. Operation

4.1 Operation process instruction

4.2 Operation

1) Illumination

- Confirm the voltage is available, then take power

- Turn on power switch B11

- Turn B9 potentiometer to adjust light

2) Sample placing

Put the sample (cover is up) on stage A6, and fix it with clamp A7. Turn B7,B8 stage knobs to put the sample into the light path. (Fig.4)

- Two samples can be placed on the stage together

3) Focusing (Fig.5)

Put 10X objective into light path, turn coarse focusing knob B13, observe from right eyepiece with right eye, then turn slowly fine focusing knob B14 to make the image clear after finding image.

- The tension adjustment ring B12 can adjust the tension of the coarse and fine focusing unit to prevent the stage from sliding down automatically and to improve the comfort of operation. Rotation clockwise makes tension decrease, and by contraries, rotation counterclockwise makes tension increase.

- Don’t turn left and right coarse and fine focusing knobs with different direction with power at the same time, if so, the focusing system will be damaged.(Fig.6)

4) Interpupillary distance and diopter adjustment

- Interpupillar distance adjustment

Interpupillar distance is different for everyone, so interpupillary distance should be adjusted before using binocular microscope. Please take the two eyepiece tubes to turn until the bright rings observed by two eyes are in superposition entirely. (Fig.7)

- Diopter adjustment (Fig.8)

As focusing for binocular, user should observe right eyepiece with right eye, and make the right eyepiece clear by focusing adjustment, then observe the left eyepiece, at the same time, adjust the diopter ring of the left eyepiece tube to make the image of left eyepiece clear as same as the right eyepiece.

5) Condenser and aperture diaphragm adjustment

- Condenser up-down

Turn the condenser focusing screw B10 to adjust the distance between condenser front and the specimen, and change the equal illumination to obtain the best brightness.

- Aperture diaphragm adjustment

Stir the handle of iris aperture of condenser to adjust size of iris aperture for adjusting the contrast of the specimen.

6) Centering field diaphragm (The process is for microscope with field diaphragm)

- Put 10X objective into light path, turn B5 field diaphragm revolver anti-clockwisely, then close field diaphragm smaller.

- Observe through eyepiece, turn B10 condenser focusing screw, adjust condenser up or down to get the image of field diaphragm

- Turn B4 condenser adjusting screw to move the image to the center of field.

- Put 40X objective into light path, turn B5 field diaphragm revolver to make the field diaphragm image bigger than field diaphragm. Please center it again if the image isn’t in center.

4.3 Usage of immersion oil objective

The 100X objective of B series microscope can be used for observation even if without immersion oil. However, adding moderate immersion oil between the front lens of 100X objective and the cover glass of specimen can make the image clearer. Please pay attention that air bubble and impurity can’t be in the immersion oil, otherwise, the image would be affected.

First of all, take 40X objective which has been focused well out from bright path, then take 100X objective into bright path. At this time, nosepiece or stage moving knob should be turned slightly, and also slightly turn fine focusing knob to clear away the air bubble of immersion oil, otherwise, the air bubble would affect the image badly.

- After immersion oil used, the oil of specimen and the microscope surface should be soon cleaned by absorbent cotton, lens paper, gauze or soft cotton cloth with moderate mixture of pure industrial alcohol and ether (proportion 1:4)

- Standard thickness 0.17mm cover glass should be chosen when high times objective used, and thickness error should be within 0.01mm, otherwise, image definition would be affected.

4.4  Points of attention after using

4.4.1 Power of microscope should be turned off after using, and the plug should be pulled down. If immersion oil used, please clean objective and specimen soon. Finally, cover microscope with dust cover.

4.4.2 Please take eyepiece and objectives out from microscope if user will stop using it for a long time, and place eyepiece and objectives into drier with drying agent. Cover microscope with dust cover.

5.Installing and using of camera and CCD attachment

5.1 Installing

Connect the C-mount with CCD camera or connect camera with camera adaptor, then connect it with c-mount, finally put it into microscope.

5.2 Using

First get a clear image from eyepiece, then put out lever on the side of trinocular head, and collect image with camera. Clear image should be in screen. Adjust B14 fine focusing knobs to get it clear if image isn’t clear.

6. Maintenance

6.1 Clean microscope

6.1.1 Don’t touch the lens with hand, Dust on lens should be cleaned by soft brush or absorbent cotton or cleaned by absorbent cotton, lens paper with the mixture of alcohol and ether (proportion 1:4).

6.1.2 Alcohol and ether all are burnt early, please take them away from fire. Be careful for turn on and off power.

6.1.3 Don’t clean painted metal and galvanizing metal with organic solvent such as alcohol, ether or the mixture of the both. Silicon cloth or soft cleaning preparation is suggested to clean it.

6.1.4 Plastic should be cleaned by soft cloth with clear water.

6.2 Environment of using and placing

6.2.1 Microscope should be used and placed in a cool, dry, non-dust, non-shake and non-corrosive gases environment.

6.2.2 Microscope should be used in environment of indoor temperature 0°-40°C and maximum relative humidity 85%.

6.2.3 Removing equipment is suggested to be installed when microscope used in heavy humidity area to avoid fungus and mist damage instrument.

6.2.4 Please pay attention to prevent microscope from violent shake and vibration in application and in carrying. Don’t drag it on the surface of worktable to avoid damage to microscope and worktable.

6.3 Replacement of bulb

6.3.1 Turn off power, and pull out plug.

6.3.2 Wait the bulb become cool.

▲ Please be sure that the bulb is cool, then follow by the nest operations.

6.3.3 Lay aside the microscope reliably, unscrew the knurled thumb screw of the lamp housing cover on the underside of base.

6.3.4 Pull over the lamp housing cover.

6.3.5 Pull out the bulb should be replaced, hold a new bulb with silk cloth to avoid fingerprint and dust affect bulb brightness and service life, and insert fully the contact pins into the bulb socket.

6.3.6 Close the lamp housing cover, and screw the knurled thumb screw.

▲After working for above 10 hours continuously, better cut off the microscope about 30 minutes.

6.4 Replacement of fuse (Shown in Fig.8)

6.4.1 Cut off power of microscope, and pull out the plug.

6.4.2 Unscrew fuse cap in the back of base, take out old fuse.

6.4.3 Replace a new fuse, then screw the fuse cap.

7.Troubleshooting

In the period of using BK series microscope, if there is any trouble occurs, please referring to the following sheet listed some common troubleshooting resolve them.

| Trouble | Causation | Remedy |
| --- | --- | --- |
| Switch on but bulb dark | Plug is unreliable | Plug in again |
|  | Bulb is broken | Change bulb |
|  | Fuse is broken | Change fuse |
| Bulb is flickering or brightness is unsteady | Bulb is unstable | Insert it again |
|  | Bulb is broken | Replacing bulb |
| Brightness of view field isn’t enough or is Uneven | Bulb specification doesn’t meet the requirement | Replacing bulb |
|  | Brightness isn’t adjusted correctly | Adjust rotation potentiometer |
|  | Objective isn’t in correct position | Make the objective in correct position |
|  | The size of iris aperture is too small | Adjust the size of iris aperture |
| Brightness of view field isn’t enough or is Uneven | Lens (objective,eyepiece, condenser, light collector) has dust | Clean it |
|  | Position of condenser is too low | Higher condenser |
| Image isn’t clear (contrast or definition isn’t enough) | Cover glass of specimen doesn’t meet the requirement | Use required thickness cover glass (0.17mm) |
|  | Cover glass of specimen isn’t in up direction | Place specimen correctly |
|  | Surface of objective lens isdirty (especially it is easy for the front lens of 40X objective to dip in immersion oil) | Clean it |
|  | Immersion oil isn’t used for 100X objective (oil) | Use immersion oil |
|  | Immersion oil doesn’t meet the requirement | Use immersion oil supplied by us |
|  | There is bubble in immersion oil | Clear the bubble way |
|  | Size of iris aperture isn’t proper | Adjust the size of iris aperture |
|  | Position of condenser is too low | Readjust the position of condenser |
| One side of image is dark or image is moving as focusing | Objective isn’t in correct position | Make the objective in correct position |
|  | Specimen isn’t placed correctly | Place specimen levelly on stage and clip it with clamp |
| Objective touches specimen as changing low times objective to high times objective | Cover glass of specimen isn’t in up direction | Place specimen correctly |
|  | Cover glass doesn’t meet the requirement | Use required thickness cover glass (0.17mm) |
| Image observed by two eyes aren’t in superposition entirely. | Interpupilary distance isn’t adjusted correctly | Adjust interpupilary distance according to two eyes |
| It is easy for eyes to be tired during observing | Diopter isn’t adjusted correctly | Readjust diopter |

8.BK Series Biological Microscope Outfits

| Items | Specification | BK200 | BK300 | BK4000 | BK5000 | Order No. |
| --- | --- | --- | --- | --- | --- | --- |
| Eyepiece | WF10×-18mm | ●● | ○○ | ●● | ○○ | WF10-18 |
|  | WF10×-20mm | ○○ | ●● | ○○ | ●● | WF10-20 |
|  | WF16×-13mm | ○○ | ○○ | ○○ | ○○ | WF16-13 |
|  | WF10×-18mm (With Pointer) | ○ | ○ | ○ | ○ | WF10P |
|  | WF10×-18mm (Reticule 0.1mm) | ○ | ○ | ○ | ○ | WF10R |
| Achromatic Objective | 4X/0.10 | ● |  |  |  | B-A4 |
|  | 10X/0.25 | ● |  |  |  | B-A10 |
|  | 40X/0.65 (S) | ● |  |  |  | B-A40 |
|  | 100X/1.25 (Oil) (S) | ● |  |  |  | B-AO100 |
| Infinity Achromatic Objective | IOS 4× |  |  | ● |  | B-Ai4 |
|  | IOS 10× |  |  | ● |  | B-Ai10 |
|  | IOS 40× |  |  | ● |  | B-Ai40 |
|  | IOS 100× (Oil) (S) |  |  | ● |  | B-Ai100 |
| Plan Objective | 4×/0.10 |  | ● |  |  | B-P4 |
|  | 10×/0.25 |  | ● |  |  | B-P10 |
|  | 20×/0.40 (S) |  | ○ |  |  | B-P20 |
|  | 40×/0.65 (S) |  | ● |  |  | B-P40 |
|  | 100×/1.25 (Oil) (S) |  | ● |  |  | B-PO100 |
| Infinity Plan Objectives | IOS 4× |  |  |  | ● | B-Pi4 |
|  | IOS 10× |  |  |  | ● | B-Pi10 |
|  | IOS 20× (S) |  |  |  | ● | B-Pi20 |
|  | IOS 40× (S) |  |  |  | ● | B-Pi40 |
|  | IOS 100× (Oil) (S) |  |  |  | ● | B-Pi100 |
| Seidentopf Binocular Head | Inclined 30°, Rotatable 360°, Interpupilary Distance 55-75mm | ● | ● |  |  | Bi-k |
| Seidentopf Binocular Head (For Infinity System) | Inclined 30°, Rotatable 360°, Interpupilary Distance 55-75mm |  |  | ● | ● | Bi-ki |
| Trinocular Head | Inclined 30°, Rotatable 360°, Interpupilary Distance 55-75mm, Light Distribution: 0:100 | ○ | ○ |  |  | Tr-k |
| Trinocular Head(For Infinity System) | Inclined 30°, Rotatable 360°, Interpupilary Distance 55-75mm, Light Distribution: 0:100 |  |  | ○ | ○ | Tr-ki |
| Nosepiece | Quadplex | ● | ● | ● | ○ | N-k4 |
|  | Quintuple | ○ | ○ |  | ● | N-k5 |
| Mechanical Stage | Travel: 50mm × 70mm; Coaxial Coarse and Fine Focusing Knobs | ● | ● | ● | ● | S-Kma |
| Condenser | Abbe N.A. 1.25 (Iris Diaphragm) | ● | ● | ● | ● | Kca |
| Koehler Illumination | 6V/20W Halogen Lamp | ● | ● | ● | ● | EK20 |
|  | 6V/30W Halogen Lamp | ○ | ○ | ○ | ○ | EK30 |
|  | LED Illumination Systems | ○ | ○ | ○ | ○ | LED |
| Field Diaphragm |  | ○ | ● | ○ | ● | K-F |
| Filter | Blue | ● | ● | ● | ● | F-BL |
| Fluorescence, Phase Contrast, Darkfield, Polarize And Incident Accessories Are Optional |

Note: “●”In Table Is Standard Attachment. “○” Is Optional Accessories

Packing Size: 540mm× 285mm×385mm     Gross Weight: 12 Kgs       Net Weight: 10 Kgs

# Text boxes

Filter

Condenser N.A.1.25

objective

Notch

Trinocular head

Binocular head

1×CCD Adaptor

Eyepiece WF10

WF10XWF10×

8

7

5

6

4

3

2

B Lamp plate screw

A Lamp plate screw

1

A2.Seidentopf binocular head

A3. Eyepiece

B4.Condenser adjusting screw (left & right)

Fig.1-2

B9. Potentiometer

A11.Condenser holder

B6. Condenser

thumb screw

B7. B8. Stage moving knobs

B5. Field diaphragm revolver

A10. Filter

A4. Nosepiece

A8. Condenser

A6. Mechanical Stage

B3. Aperture diaphragm lever

A7. Clamp

A9. Notch for attachment

B2. Diopter adjuster

B1. Head thumb screw

A5. Objective

A1.Body

B14. Fine focusing knobs

B13. Coarse focusing knobs

B12. Tension adjustment ring

B11.Power Switch

A13. Power Input

A14. Fuse

B10.Condenser focusing screw

A12. Handle

Fig.1-1

##### Fig.2

Fig.3

Adjust field diaphragm (B5 field diaphragm revolver), adjust center screw B4 to center field diaphragm

Adjust A8 condenser (B10. condenser focusing screw) and aperture (B3. Aperture diaphragm lever)

adjust A2 head to make two fields into one; adjust B2 diopter adjuster to make diopter for two eyes same.

urn A4 nosepiece, put 10X objective into light path, turn B13,B14 coarse and fine focus knobs

Put the sample on A6 stage,(open the clamp, and put the sample in it), turn  B7,B8 stage knobs to put the sample into light path

Turn on B11 power switch, and turn B9 potentiometer

Observing

Put the objective into light path, and focus

6) Adjust and center field diaphragm

5) Adjust condenser aperture diaphragm

4) Adjust interpupillary distance and diopter

3) Put 10X objective into light path, and focus the sample

2) Put the sample

1) Turn power switch, adjust light

Fig.4

Fig.5

Fig.6

Fig.7

Fig.8